## Sample Experiment Log
astraX — Personal Experiment Log
Owner: ___________________ Track(s): ___________________ Updated: ___________________
Hypothesis format: 'If we do X, then Y will happen, because Z.' Success Criteria: measurable, written BEFORE the experiment. Status: Not Started ¦ In Progress ¦ Complete ¦ Blocked ¦ Inconclusive
Track | Experiment Title | Description / Hypothesis | Experiment Design | Success Criteria | Target End Date | Dependencies | Outcome / Finding | Next Action | Status
T1 AstraX Device | GEOMINFO_ts sensitivity on Mo | If we reduce GEOMINFO_ts by 20%, Mo overcall on 316 SS will drop below 5% delta | Run 5 spectra on NIST SRM 361 (316 SS). Vary GEOMINFO_ts from 0.8x to 1.2x in steps. Log Mo% at each setting. | Mo delta vs certified < 5% on ≥3 of 5 runs | 2026-06-20 | HIL jig live (T1-Jul); DPP gain locked |  |  | In Progress
T2 AstraX Customer Cloud | ScrapSense onboarding flow — scrap yard pilot | If we reduce onboarding steps from 7 to 4, pilot user activation time will drop below 10 minutes | A/B test: current 7-step flow vs simplified 4-step flow with 2 scrap yard users. Measure time-to-first-measurement. | Median activation time < 10 min; user rates flow ≥ 4/5 | 2026-07-15 | ScrapSense MVP spec finalised (T2-Jul) |  |  | Not Started
T3 AstraX Ops Cloud | Implement sales agent | Create a Agent to support astraX sales workflows online, Amazon like experience | Refer to HW companies who already implemented this capability online | Significantly removes grunt work involved with sales operations, customer is delighted
T4 Manufacturing partners | India CM #1 — PCB assembly cost vs US baseline | India CM can produce 3-board PCB assembly set at ≤60% of current US quote ($900 target vs $1,500 US) | Send Gerbers + BOM to top 3 India CM candidates. Request unit price at 10-unit and 50-unit MOQ. Compare to US baseline quote. | At least 1 CM quotes ≤$900 at 50-unit MOQ with ≥IPC Class 2 quality standard | 2026-07-10 | BOM India-sourceable audit complete (T4-Jun); NDA signed |  |  | Not Started
T5 Business | Break-even sensitivity — COGS reduction impact | If India CM hits $9K COGS target (vs $10.66K), break-even drops from 38 to ≤28 units/yr — within 2027 target | Update Device Economics sheet with India CM quotes as received. Re-run sensitivity table. Plot break-even vs COGS mult. | Break-even ≤30 units/yr confirmed at India CM price point | 2026-08-15 | T4 CM RFQ responses received (T4-Jul); T2 pricing confirmed |  |  | Not Started
T6 Sales partner

## All Experiment Summary
astraX — Team Experiment Summary
HOW TO USE: Each team member updates their Experiment Log regularly. Can Filter by Track or Status to run standup. Stats panel auto-counts from col J.
Owner | Track | Experiment Title | Function | Parent Item | Description / Hypothesis | Experiment Design | Success Criteria | Target End Date | Dependencies | Test outcome / Finding | Next Action | Status | STATUS SUMMARY | Count
Vijay | T1 Device | astraX calibration framework |  |  | A intuitive calibration app framework for developing XRF applications, to be used by XRF application engineer | Initially build on a PC, later make this a calibrationAgent in the Cloud in future | quick development and debug of XRF applications | ? | XRS library calibration design | Ability to create calibrations |  | Not Started | Not Started | 6
Sreekar | T1-Device | Battery Management System |  |  | If we set up a evaluation board equivalent of our pcb layout, we can confirm that our design is correct | Get EV2500 programmer, set up individual boards, connect real battery, connect stm for monitoring | Able to properly charge the battery, use the battery, and view the current status of it. | 2026-06-14 |  |  | Wait for EV2500 and battery | In Progress | In Progress | 4
Sreekar | T1-Device | Set up STM base firmware | [10 empty] | Complete | 0
Gopinath | T1-Device | Calculate weight for the device from design | Mechanical |  | calculate weight from design file and compare with physical sample. work on weight reduction if the device is more than 1.5kg | build the EB1 mock up system and measure , optimize the design for EB2 to meet weight target | meet the weight target(,1.5kg) is ideal |  | EB1 mock up sample |  |  | In Progress | Blocked | 0
Gopinath | T1-Device | PMMA assembly with king tech CG | Mechanical |  | identify optimized method to integrate PMMA cover into CG with transperant glue or adhesive | [6 empty] | In Progress | Inconclusive | 0
Ravi | T1-Device | Battery design with integration of terminals | Mechanical |  | Lock/unlock of Battery with premium feel. Reliable Connection of battery terminal with handle teminals. | Develop battery housing prototype with integrated terminal design. Perform insertion/removal cycle testing, terminal alignment verification, and user evaluation for locking feel. Test minimum 20-50 insertion cycles. | 1. Battery locks securely and removes easily without excessive force. 2. No terminal damage after repeated cycles. 3. No intermittent electrical connection during handling/vibration. 4. User feedback indicates acceptable premium feel. 5. Water does not "pool" enough to short circuit + and - contacts. (See also IP54 entry) 6. Minimal freeplay when attached (rotary and XY). | 2026-10-15 | Battery housing prototype, terminal selection & samples, handle assembly prototype, ID input |  |  | Not Started | TOTAL | 10
Aditya | T1-Device | RAD LED uniform illumination | Mechanical |  | RAD LED is easily visible, with bright and uniform illumination, in daylight conditions, from all intended angles of view. | View in dark and bright sunlit conditions to evaluate. | Users say that it's easy to see if the LED is lit while viewing in daylight conditions. Users say that the LED bar appears to be uniformly lit. |  |  |  |  | Not Started
Aditya | T1-Device | Trigger Button user experience | Mechanical |  | The trigger actuation can be designed to provide a highly satisfying experience without fatigue and exuding a sense of high quality. | Design drop-in module that fits where E-Switch trigger was mounted. Make it adjustable in spring force and travel. Do we need a detent or is simple linear travel enough with sound and LED light activation? | Users say that the trigger feels of high quality, does not cause fatigue after 50 actuations. |  |  |  |  | Not Started
Gopinath | T1-Device | Power button light illumination and tactile feel | Mechanical |  | Power button illumination is easily visible, with bright and uniform illumination, in daylight conditions, from all intended angles of view. | View in dark and bright sunlit conditions to evaluate. | Users say that it's easy to see if the PWR button is lit while viewing in daylight conditions. Users say that the PWR button appears to be uniformly lit. |  |  |  |  | Not Started
Aditya | T1-Device | Type C connector life cycle test | Mechanical |  | meet cycle test 5000 to 10000 insertion and extraction cycle for type C | [6 empty] | Not Started
Ravi | T1-Device | IP54 Sealing | Mechanical |  | EB2 prototypes exposed to IP54 water and dust exposure will not allow ingress. | Design water-resistant foam seals or O-rings at all joints that have been analyzed to insure enough compression to resist water/dust and not have excessive force that might distort parts and cause cosmetic misalignments. | EB2 prototypes exposed to IP54 water and dust exposure will not allow ingress. |  |  |  |  | In Progress
|  | ID - Final Form | Mechanical |  | Vijay cited some updates needed on the form, e.g. handle size, battery width, rounding the top cover, making bezel narrower (side view). We're deciding what to require in EB1 and what can be pushed into EB2. Final ID, Fresco adjustments required early/during EB2 design cycle.
|  | Handle, grip size | Mechanical |  | Increase grip size for better comfort/control. Work with ID and Vijay to define the size and contour that works well for 25th to 75th percentile hand size. With gloves?
|  | UI Controls | Mechanical |  | Redesign of Pwr, Rocker is expected during EB2. Final ID, Fresco revisions are needed early/during EB2 design cycle.
|  | Macro Camera | Mechanical
|  | Kapton Window | Mechanical |  | Custom die-cut, peel-and-stick window
|  | Injection Molding | Mechanical |  | 2-shot or insert-molded Handle L/R. Drafts and Split-lines are modeled. Need ID texture specs to define draft angles.
|  | CMF | Mechanical |  | Final ID, Fresco adjustments required. Fresco work must be done early in EB2 design cycle. Define CMF citing standard color and finish standards by our supplier.
|  | Drop Survival | Mechanical |  | EB1 must incorporate a design strategy that is intended to pass 36" drop. This prototype need not pass. We should run tests on a mock-up with housing and frame only. EB2 must pass Drop Requirement, because high-confidence is needed before investing in PB tooling. Most fragile components: >xray tube >detector >display
Note on A11: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Single owner for the experiment, others can support @Aditya Reply: i have changed the ownership to aditya

## How To Use
astraX Experiment Log — How To Use
COLUMN | GUIDANCE
Track | Select from dropdown: T1 Device through T6 Sales. Links your experiment to the Jun–Nov roadmap.
Experiment Title | Short scannable name. Use in standups. Example: 'GEOMINFO_ts sensitivity on Mo'. Any atomic experiment is NOT bigger than 2 weeks.
Description / Hypothesis | Write as: 'If we do X, then Y will happen, because Z.' Be specific. Bad: 'Test the quant model.' Good: 'If we reduce GEOMINFO_ts by 20%, Mo overcall on 316 SS will drop below 5% delta.'
Experiment Design | How will you run it? List: materials/samples needed, steps, number of runs, measurement method. Enough detail that someone else could run it in your absence.
Success Criteria | WRITE THIS BEFORE YOU START. What specifically does 'pass' look like? Must be measurable. Example: 'Mo delta vs certified < 5% on ≥3 of 5 runs.' Prevents moving goalposts after the fact.
Target End Date | Pick a realistic date. If it slips, update it and note why in Dependencies.
Dependencies | What must be true before this experiment can start or finish? Format: 'HIL jig live (T1-Jul)' or 'India CM RFQ received (T4-Jul)'. Surface blockers during standup
Outcome / Finding | Fill in AFTER the experiment. State the actual result in one sentence. Then note what it means: 'Pass — confirms hypothesis' or 'Fail — Mo delta was 9%, not <5%. Energy calibration may be the root cause.'
Next Action | What does this result trigger? Must be actionable. Examples: 'Run experiment #2 with adjusted detector angle' or 'Update CalcEngine GEOMINFO_ts to 0.8x and retest' or 'Escalate to Vijay — blocked on DPP gain setting.'
Status | Select from Drop Down (Not Started → In Progress → Complete or Blocked or Inconclusive). Update every Friday. Row turns green on Complete, red on Blocked automatically.
WEEKLY RITUAL
weekly standup calls | Update XL before the standup
Monday (Leads meeting) | Review weekly progress and challenges, prioritize
HYPOTHESIS TIP | 'Experiments' without hypotheses are just tasks. The discipline of writing 'If X then Y because Z' forces clarity on what you're actually trying to learn — which is the whole point of experiment-based execution.